Tender submission form

small-scale Public procurement entitled

„FTK/UPOL – Isokinetický dynamometr II“

Identification details of the tenderer and Its representatives

Identification details of the tenderer

| Business name or full name: | Insert name |
| --- | --- |
| Registered office: | Insert registered office |
| Company ID no.: | Insert no. |
| VAT ID no.: | Insert no. or information that the tenderer is not a VAT payer |
| Entry in CR: | Insert court, section and file no. |
| Bank details: | Insert bank name |
| Bank account no.: | Insert bank no. |
| The tenderer is a small or medium-sized enterprise : |  |
| Legal representative: | Insert name, surname and position of the representative who will sign the Contract in the event of selection |
| Joint participation of tenderers in the selection procedure: | Business name/full name, registered office, company ID no., VAT no. |
|  | Business name/full name, registered office, company ID no., VAT no. |
|  | Business name/full name, registered office, company ID no., VAT no. |

Identification details of tenderer’s representatives

| Tenderer’s contact person regarding the tender |  |
| --- | --- |
| Name and surname: | Insert name and surname |
| Telephone: | Insert telephone no. including international dialing code |
| E-mail: | Insert e-mail adress |
| Tenderer’s person authorized to act in technical matters during performance of the Contract |  |
| Name and surname: | Insert name and surname |
| Telephone: | Insert telephone no. including international dialing code |
| E-mail: | Insert e-mail adress |

Tenderer’s declarations

The tenderer hereby declares on its honor

with regard to the tender conditions, that it

has carefully reviewed the tender conditions of the Public procurement, including any clarifications, supplements and amendments thereto, has understood them and agrees to them,

accepts E-ZAK as the exclusive means of communication in the selection procedure of the Public procurement,

acknowledges that, in communication carried out via E-ZAK, a data message is deemed delivered upon its receipt at the electronic address of its addressee.

with regard to the binding commercial terms of the Public procurement, that it accepts them in the wording provided by the Contracting authority;

with regard to the responsible performance of the Public procurement, that

in performing the subject of the Public procurement, it shall ensure, in respect of all employees participating in the performance of the Public procurement, compliance with labor law regulations, employment regulations and occupational health and safety regulations;

it shall impose the same obligations on its subcontractors directly participating in the performance of the Public Procurement (excluding suppliers of materials or goods in the case of resale);

with regard to conflict of interest within the meaning of the Subsidy rules, that no conditions for the existence of a conflict of interest within the meaning of Section 44 ZZVZ are met; in particular, it is not directly or indirectly influenced in the selection procedure in relation to the Contracting authority or to entities involved in the preparation of the tender conditions, nor does it have any special relationship with such persons (e.g. ownership, personnel or similar);

with regard to conflict of interest within the meaning of Section 4b ZOSV, that:

it is not a commercial company in which a public official referred to in Section 2(1)(c) ZOSV or a person controlled by such official, holds a share representing at least 25% participation in that commercial company; and

if it demonstrates fulfilment of any part of qualification through another person (subcontractor), such other person is not a commercial company in which a public official referred to in Section 2(1)(c) ZOSV or a person controlled by such official, holds a share representing at least 25% participation in that commercial company;

with regard to international sanctions, that

it is familiar with the applicable international sanctions, in particular Regulation (EU) No. 833/2014 concerning restrictive measures in view of Russia’s actions destabilising the situation in Ukraine, as amended by Council Regulation (EU) No. 2022/1269 of 21 July 2022;

it does not offer performance subject to applicable international sanctions;

if a contract for the Public procurement is concluded with it, it shall ensure for the entire duration of its performance that

the above shall also apply to all its subcontractors used in the performance of the Public procurement whose share in the subject of the Public procurement exceeds 10%;

it shall immediately replace any subcontractor meeting the above criteria if international sanctions are imposed on such subcontractor during the performance of the Public procurement.

Qualification

Basic eligibility

The tenderer hereby declares on its honor that it is eligible to perform the Public procurement within the meaning of Section 74 ZZVZ.

Professional eligibility

The tenderer hereby declares on its honor that it fulfils the requirement of Section 77(1) ZZVZ. This fact shall be demonstrated by a sworn declaration.

Technical qualification

The Contracting Authority does not require demonstration of technical qualification.

Evaluated criteria

The tenderer offers a total tender price Insert ammount CZK excluding VAT.

Technical specification

| Item no. | Basic technical parameters | Required level of technical parameters | Offered level of technical parameters |
| --- | --- | --- | --- |
| 1. | Isokinetic dynamometer | Insert manufacturer name and model designation |  |
| 1a | Number of units | 1 unit |  |
| Functional requirements |  |  |  |
| 1b | Testing in concentric, eccentric (reciprocal modes concentric-concentric, concentric-eccentric, eccentric-concentric and eccentric-eccentric) and isometric muscle work modes. | YES |  |
| 1c | Controlled passive motion (CPM) with the possibility of reversing direction after exceeding an adjustable eccentric torque threshold. | YES |  |
| 1d | Dynamically controlled velocity profiles with the following individual parameters (separately for both directions of movement): • Acceleration and deceleration at least 8,000 °/s². • Maximum angular velocity up to 550°/s. • Maximum torque ≥ 750 Nm in both concentric and eccentric phases. • Initiation of movement upon reaching an adjustable torque threshold, including the possibility of expression as % of maximum isometric test. | YES |  |
| Compensation and device settings |  |  |  |
| 1e | Gravity compensation with automatic storage for the tested individual, joint and plane of movement. | YES |  |
| 1f | Reproducible fixation of the rotational position of the dynamometer relative to the seat/bench (resolution min. 0.5°). | YES |  |
| 1g | Stability of the dynamometer position – resistance to maximum torque in all fixed positions | YES |  |
| 1h | Minimum height of dynamometer axis 60 cm. | Min. 60 cm. | Insert offered parameter level |
| 1i | Vertical backrest adjustment at least 24.5 cm. | Min. 24,5 cm. | Insert offered parameter level |
| 1j | Backrest tilt 85°–0°, with mechanical stabilization lock. | YES |  |
| 1k | Adjustable seat tilt with possibility of horizontal (0°) setting for testing in supine position. | YES |  |
| Memory and automatic settings |  |  |  |
| 1l | Sensor-based storage of seat and dynamometer position settings according to tested person and plane of movement (height, tilt, axis position, backrest tilt and shift). | YES |  |
| 1m | Automatic adjustment of seat and dynamometer position according to individually stored data (memory function). | YES |  |
| Upper body fixation |  |  |  |
| 1n | Fixation in vertical and horizontal planes, with integrated shoulder pads adjustable in height, width and depth; shoulder pads cannot be replaced by separate straps. | YES |  |
| Monitor Requirements |  |  |  |
| 1o | Monitor on a movable arm with reach of several metres for ergonomic adjustment relative to the tested person and operator. | YES |  |
| 1p | No loose cables in the working area (prevention of tripping). | YES |  |
| 1q | In supine position, possibility to place the monitor above the tested person’s body for natural field of vision. | YES |  |
| 1r | Height adjustable by at least 40 cm for correct monitor alignment relative to the eyes of the tested person and operator. | YES |  |
| Alignment and stabilization |  |  |  |
| 1s | Laser alignment of the dynamometer axis to the joint axis of the tested person for maximum accuracy. | YES |  |
| 1t | Sliding system with possibility of attaching adapters on both sides of the seat/bench – enables individual stabilisation and joint isolation. | YES |  |
| Software, data output and visualization |  |  |  |
| 1u | Software enabling recording of measurements, basic data processing and export. | YES |  |
| 1v | Export of raw data with frequency min. 200 Hz (5 ms) including measured values, test data and settings. | YES |  |
| 1w | Analog output ±5 V for simultaneous recording of torque and angular position during movement. | YES |  |
| 1x | Real-time visualisation in graphical interface – display of torque–angle curveas biofeedback. | YES |  |
| 1y | Possibility to store and display reference curves (scalable, for specific exercise and person) and overlay them during measurement. | YES |  |
| Module compatibility |  |  |  |
| 1z | Torsionally rigid coupling for connection and sharing of measuring/drive unit with additional modules (rotational or linear – leg press). | YES |  |

Dodavatel vyplní „ANO“ pouze v případě, že je malým nebo středním podnikem ve smyslu doporučení Komise 2003/361/ES o definici mikropodniků, malých a středních podniků. Bližší informace včetně českého znění uvedeného doporučení jsou uvedeny v Uživatelské příručce k definici malých a středních podniků, která je dostupná na adrese https://publications.europa.eu/cs.

V případě společné účasti dodavatelů odkliknout zaškrtávací pole. Veškerá prohlášení ve Formuláři nabídky činí v případě společné účasti dodavatel s uvedeným zástupcem i za ostatní dodavatele.